TEST REPORT

Page  1  of  3

REPORT NO:					TEST DATE:			SERIAL NO:

EQUIPMENT CLASSIFICATION:	Type A2 Biological Safety Cabinet

TEST REQUIREMENTS:	NSF / ANSI 49 – 2002 Annex A				MODEL NO:

MANUFACTURER:								LOCATION:

CLIENT:										PHONE NO:

ADDRESS:									CONTACT:

| TEST METHOD | TEST DESCRIPTION | REQUIREMENT | READING | RESULT |
| --- | --- | --- | --- | --- |
| Method 9 | Average Downflow Velocity: Uniformity: | Average = 0.40 to 0.45 m/s Uniformity : +20% (0.08m/s Max) -20% (0.08m/s Max) | m/s + % - % |  |
|  | Pressure Differentials: | N/A | LAF Pa |  |
| Method 3 | HEPA Filter Integrity | <0.01% DOP Penetration: | LAF <0.01% EXH <0.01% |  |
| Method 11 | Airflow Smoke Patterns Method C1 Method C2 Method C3 | Smooth Flow Smooth Flow No Smoke Escape or Inflow | N/A N/A N/A |  |
| Method 10 | Inflow Air Velocity % Exhaust Air | 0.5m/sec Minimum 0.03m³/sec per 0.3m of work area width minimum | m/sec m³/sec per 0.3m width minm |  |
| Method 5 | Lighting Intensity: | Minimum Average ≥650 lux: Minimum Reading ≥430 lux: | lux lux |  |
| Method 4 | Noise Level: | Maximum 67dB(A) | dB(A) |  |

This Laboratory is accredited by the National Association of Testing Authorities, Australia. Accreditation No:  14953. The tests reported herein have been performed in accordance with its scope of accreditation. This document shall not be reproduced, except in full.

TESTED BY:				N.A.T.A. SIGNATORY:					DATE:

NAME (PRINT):