TEST REPORT

Page  1  of  3

REPORT NO:					TEST DATE:				SERIAL NO:

EQUIPMENT CLASSIFICATION:	Cytotoxic Drug Safety Cabinet

TEST REQUIREMENTS:	AS2567							MODEL NO:

MANUFACTURER:									LOCATION:

CLIENT:											PHONE NO:

ADDRESS:										CONTACT:

| TEST METHOD | TEST DESCRIPTION | REQUIREMENT | READING | RESULT |
| --- | --- | --- | --- | --- |
| AS1807.1 | Average Air Velocity and Uniformity | Average = 0.40 to 0.50 m/s Uniformity +15% -15% | m/s + % - % |  |
|  | Pressure Differentials: | N/A N/A | LAF Pa EXH Pa | N/A N/A |
| AS1807.5 | Work Zone Integrity | <0.01% DOP Penetration: | <0.01% |  |
| AS1807.6 | HEPA Filter Integrity | <0.01% DOP Penetration: | <0.01% |  |
| AS1807.15 | Lighting Intensity: | Minimum 800 lux | Lux |  |
| AS1807.20 | Sound Level: | Maximum 62 dB(A) | dB(A) |  |
| AS1807.22 | Barrier Containment: | <0.01% per 10Seconds | <0.01% |  |
| AS1807.23 | U.V. Intensity: | Minimum 400mW/m² | mW/m² |  |

This Laboratory is accredited by the National Association of Testing Authorities, Australia. Accreditation No:  14953. The tests reported herein have been performed in accordance with its scope of accreditation. This document shall not be reproduced, except in full.

TESTED BY:				N.A.T.A. SIGNATORY:					DATE:

NAME (PRINT):